## Slide 1
GlucoTwin
Predictive Metabolic Digital Twin for Hypoglycemia Risk
Ronak Jain · Aishni Rathore
Jaypee Institute of Information Technology, Noida-62
CRP Digital Twin Challenge 2026 · Happiest Health

## Slide 2
Problem Statement
• Hypoglycemia (≤ 70 mg/dL) is the most critical acute hazard for Type 1 Diabetes patients • Risk is highest during sleep — when symptom perception is absent • Current CGM threshold alerts: 29.21 raw false alarms/day causing severe alert fatigue • Static alarms ignore physiological context: IOB, COB, circadian vulnerability • No counterfactual simulation: cannot ask 'what if the patient takes 15g carbs?' • No continuous patient memory: every moment evaluated in isolation
2

## Slide 3
Healthcare Use Case
• Target Users: Type 1 Diabetes patients on CSII/MDI insulin therapy with CGM sensors • Clinical Setting: Nocturnal hypoglycemia prevention and early-warning decision support • Clinician Decision Support: Calibrated 60-min risk alerts before threshold crossing • Patient Empowerment: Interactive What-If simulator for rescue carb intervention planning • Hospital/Clinic Use: Multi-patient dashboard with per-patient personalized twin states • Regulatory Scope: Research prototype — clinician decision support only (not FDA-cleared)
3

## Slide 4
Technical Stack
ML & Backend
Frontend & Infrastructure
• Python 3.10/3.11 • LightGBM + RandomForest ensemble • SHAP (explainability) • Scikit-learn (isotonic calibration) • NumPy / Pandas / SciPy • Joblib (model serialization) • pytest (38 tests, CI verified)
• Streamlit (clinical dashboard) • Plotly (interactive charts) • GitHub Actions CI/CD • Python 3.10 + 3.11 CI matrix • edge-tts / moviepy (demo video) • python-pptx / reportlab (docs) • Dataset: HUPA-UCM (Mendeley DOI)
4

## Slide 5
AI/ML Model & Framework Details
• Primary Model: LightGBM 60-min hypoglycemia classifier (50 features) • Validation: Leave-One-Patient-Out (LOPO) on 24 patients — ROC-AUC 0.970 ± 0.027 • Calibration: Isotonic regression — Brier score 0.069 → 0.039 (−43.5%), ECE −94.5% • Explainability: SHAP TreeExplainer — real-time top-K risk factor attribution • Trajectory: Damped-velocity ODE (k=0.015 min⁻¹) — TTT MAE 32.55 min over 11,471 pairs • Alert FSM: Hysteresis state machine — 4.39 episodes/day vs 29.21 raw point alarms • OOD Detection: Mahalanobis distance + CGM continuity → twin_confidence 0.0–1.0
5

## Slide 6
5 Core Digital Twin Differentiators
• ① Continuous Patient Memory — persistent episode log with strict backward temporal guards • ② Trajectory Modeling with Validated Uncertainty — damped ODE + empirical MAE tiers • ③ Counterfactual What-If Simulation — in-silico rescue carb & bolus scenario engine • ④ Stateful Alert Engine with Hysteresis — 4.39 episodes/day (vs 29.21 raw false alarms) • ⑤ Model Self-Awareness & OOD Guardrails — TwinState v6 exposes full empirical metrics
6

## Slide 7
Authoritative Empirical Results
• LOPO ROC-AUC: 0.970 ± 0.027 | PR-AUC: 0.858 ± 0.099 (24 patients, patient-independent) • LOPO Sensitivity: 0.958 ± 0.041 | Specificity: 0.821 ± 0.088 | F2: 0.754 ± 0.065 • TTT MAE: 32.55 min | Median AE: 17.60 min | Within 15 min: 46.0% (11,471 pairs) • Brier Score: 0.0690 (raw) → 0.0390 (calibrated) | ECE: 0.1006 → 0.0055 • Full-Cohort Ablation AUC: 0.9468 | PR-AUC: 0.8132 | Sensitivity: 0.8628 • Episode Detection: 300/300 (100%) | Mean Lead Time: 47.4 min • Alert Episodes: 4.39/day (vs naive 29.21 raw point-level alarms/day)
7

## Slide 8
Streamlit Clinical Dashboard (app.py)
• Live Inference Mode: real-time risk scoring from production LightGBM + IsotonicCalibrator • Patient Navigator: Sidebar dropdown for all 25 HUPA-UCM cohort patients • Current State Panel: glucose, 15-min velocity, risk score, alert status, trend • What-If Simulator: carb intervention sliders (0–80g), bolus, exercise, 30–180 min horizon • SHAP Explanations: top risk drivers computed live per inference call • OOD Confidence Panel: twin_confidence score + Mahalanobis distance visualization • Demo Fallback Mode: verified static payload if model artifacts are absent
8

## Slide 9
System Architecture Overview
• Data Layer: HUPA-UCM CGM + Bolus + Meal + Activity data ingestion • Processing Layer: Leakage-safe preprocessing → 50-feature engineering pipeline • Twin Core: LightGBM + Isotonic Calibrator → Alert FSM → Patient Memory → OOD • Simulation Layer: MetabolicSimulator v2 (Bergman 2-compartment ODE) • API Layer: GlucoTwinAPI → TwinState v6 JSON (full snapshot, SHAP, What-If) • Presentation Layer: Streamlit + Plotly live dashboard (app.py) • Full architecture diagram: docs/submission/GlucoTwin_Architecture_Diagram.pdf
9

## Slide 10
Limitations & Ethical Considerations
• RESEARCH PROTOTYPE — NOT A MEDICAL DEVICE (not FDA/CE cleared) • Observational cohort: 25 Spanish adults with T1D — generalizability unknown • Sensor-specific: FreeStyle Libre 2 CGM; other devices may differ in factory calibration • No Closed-Loop Control: open-loop decision support only — not integrated into AID systems • What-If simulations are mathematical projections, NOT clinical prescriptions • All predictions are statistical probabilities — never replace physician judgment
10

## Slide 11
Team GlucoTwin
CRP Digital Twin Challenge 2026 · Happiest Health
Ronak Jain (ML/Backend) · Aishni Rathore (Dashboard/Frontend)
Jaypee Institute of Information Technology, Noida-62
github.com/Ronak1167/glucotwin | CI: ✅ 38/38 tests passing